## Sheet1
CT Scanner ID: ____________________________________________
CT WATER PHANTOM DAILY QC
Week of ____________________
| RANGE | SUN | MON | TUE | WED | THU | FRI | SAT
Technogist's Initials
Water Center Mean CT # | -3.0 to +3.0
Water Center Stand Dev. | 4.0 to 5.7
Visual Artifacts f. centered phantom | Yes / No
Visual Artifacts f. uncentered phantom | Yes / No
CT WATER PHANTOM DAILY QC
Week of ____________________
| RANGE | SUN | MON | TUE | WED | THU | FRI | SAT
Technogist's Initials
Water Center Mean CT # | -3.0 to +3.0
Water Center Stand Dev. | 4.0 to 5.7
Visual Artifacts f. centered phantom | Yes / No
Visual Artifacts f. uncentered phantom | Yes / No
CT WATER PHANTOM DAILY QC
Week of ____________________
| RANGE | SUN | MON | TUE | WED | THU | FRI | SAT
Technogist's Initials
Water Center Mean CT # | -3.0 to +3.0
Water Center Stand Dev. | 4.0 to 5.7
Visual Artifacts f. centered phantom | Yes / No
Visual Artifacts f. uncentered phantom | Yes / No
CT WATER PHANTOM DAILY QC
Week of ____________________
| RANGE | SUN | MON | TUE | WED | THU | FRI | SAT
Technogist's Initials
Water Center Mean CT # | -3.0 to +3.0
Water Center Stand Dev. | 4.0 to 5.7
Visual Artifacts f. centered phantom | Yes / No
Visual Artifacts f. uncentered phantom | Yes / No
CT WATER PHANTOM DAILY QC
Week of ____________________
| RANGE | SUN | MON | TUE | WED | THU | FRI | SAT
Technogist's Initials
Water Center Mean CT # | -3.0 to +3.0
Water Center Stand Dev. | 4.0 to 5.7
Visual Artifacts f. centered phantom | Yes / No
Visual Artifacts f. uncentered phantom | Yes / No
CT WATER PHANTOM DAILY QC
Week of ____________________
| RANGE | SUN | MON | TUE | WED | THU | FRI | SAT
Technogist's Initials
Water Center Mean CT # | -3.0 to +3.0
Water Center Stand Dev. | 4.0 to 5.7
Visual Artifacts f. centered phantom | Yes / No
Visual Artifacts f. uncentered phantom | Yes / No
2024-10-18
CT Scanner ID: ____________________________________________
CT WATER PHANTOM WEEKLY QC
| RANGE | DATE | DATE | DATE | DATE | DATE | DATE | DATE
Technogist's Initials
Water Center Mean CT # | -3.0 to +3.0
Water Center Stand Dev. | 4.0 to 5.7
Visual Artifacts f. centered phantom | Yes / No
Visual Artifacts f. uncentered phantom | Yes / No
Low Contrast Resolution (3.0mm) (64 slice scanners only) | < 5.3 HU
Mean Water CT # | ----
Mean Plexiglas CT # | ----
Difference Plexiglas - Water | 108 - 132
Hi Contrast Resolution - Stand Recon (1.6 mm bar - Largest bar) | > 33.0
Hi Contrast Resolution - Stand Recon (0.8 mm bar) | > 8.9
Hi Contrast Resolution - Edge Recon (0.6 mm bar - Next to smallest bar) | > 37.0
Slice Thickness 5.00 mm | 5 ± 0.5
Slice Thickness 2.50 mm | 2.5 ± 0.5
Slice Thickness 1.25 mm | 1.25 ± 0.5
Internal Axial Laser Accuracy | ≤ 2mm
90 / 270 Laser Accuracy (up/down) | ≤ 2mm
0 Laser Accuracy (left/right) | ≤ 2mm
CT WATER PHANTOM WEEKLY QC
| RANGE | DATE | DATE | DATE | DATE | DATE | DATE | DATE
Technogist's Initials
Water Center Mean CT # | -3.0 to +3.0
Water Center Stand Dev. | 4.0 to 5.7
Visual Artifacts f. centered phantom | Yes / No
Visual Artifacts f. uncentered phantom | Yes / No
Low Contrast Resolution (3.0mm) (64 slice scanners only) | < 5.3 HU
Mean Water CT # | ----
Mean Plexiglas CT # | ----
Difference Plexiglas - Water | 108 - 132
Hi Contrast Resolution - Stand Recon (1.6 mm bar - Largest bar) | > 33.0
Hi Contrast Resolution - Stand Recon (0.8 mm bar) | > 8.9
Hi Contrast Resolution - Edge Recon (0.6 mm bar - Next to smallest bar) | > 37.0
Slice Thickness 5.00 mm | 5 ± 0.5
Slice Thickness 2.50 mm | 2.5 ± 0.5
Slice Thickness 1.25 mm | 1.25 ± 0.5
Internal Axial Laser Accuracy | ≤ 2mm
90 / 270 Laser Accuracy (up/down) | ≤ 2mm
0 Laser Accuracy (left/right) | ≤ 2mm
2024-10-18

## Sheet2
CT Scanner ID: ____________________________________________
CT MONTHLY QC
| DATE | DATE | DATE | DATE | DATE | DATE
Technogist's Initials
VISUAL CHECKLIST
GANTRY CHECKS:
Table height indicator functioning
Table position indicator functioning
Angulation indicator functioning
Laser localization light functioning
Smoothness of table motion acceptable
X-ray on indicator functioning
CONTROL CONSOLE:
Exposure switch functioning
Display window width/level
Panel switches/lights/meters working
X-ray on indicator functioning
Warning labels present
Intercom system functioning
GRAY LEVEL PERFORMANCE OF CT SCANNER ACQUISITION DISPLAY MONITOR
CHECKS ON SMPTE PATTERN:
5% patch visible within 0/5% patch
95% patch visible within 95/100% patch
All gray levels (0% to 100%) are distinct from adjacent levels
PASS = P
FAIL = F